## Slide 1: Robert Stobie Prime Focus Imaging Spectrograph Linear Spectropolarimetry - Calibration K Nordsieck
RSS polarimetric optics
what needs to be calibrated
calibration data maintained by SALT
calibration is applied in distributed software polsalt
known problems in current calibration
plans for future
Danièl: Precision Me: Accuracy

## Slide 2: RSS Spectropolarimetry - overview
RSS Spectropolarimetry: object on-axis at focal plane, though waveplate, dispersed by grating, and split into O and E spectra
Polsalt pipeline: wavecal, extract, rawstokes, calibrate, plot.
Sraw =
½ [(O1 - O2)/(O1 + O2) -
(E1 - E2)/(E1 + E2)]
beamsplitter
UV Calcite
Prism mosaic
150 mm
grating
E
O
WR 79 900 l/mm
20220923
halfwave
quarterwave
stack of 3: MgF2/crystal quartz
300μ x 100, 60 mm
Linear spectropolarimetry: images taken at 2 halfwave plate rotations 1,2 45º apart.
  Qraw : hw pos 0,4: 0º,45º
  Uraw : hw pos 2,6: 22.5º,37.5º

## Slide 3: Calibration
Note: much of this is in the Observer's Guide:
http://astronomers.salt.ac.za/wp-content/uploads/sites/71/2022/06/3170AM0013_Polarimetry_Observers_Guide_V1.2.pdf
 |  | Status |  | 
Calibration | Origin | On-Axis | Full FOV | Full Track
Linear Pol Zero-Point | Tel Mirrors. λ dep | Done | in work | in work
Linear Pol Efficiency | Waveplate. λ dep | Done | in work | in work
Linear Pol PA Zero-Point | Rho, waveplate rotator | Done | NA | NA
Circular Pol Zero-Point | ?, likely negligible | TBS | TBS | TBS
Circular Pol Efficiency | Waveplate | TBS | TBS | TBS
Linear-Circular Conversion | ADC, Waveplate | TBS | TBS | TBS
What needs calibrating? here: grating spectropolarimetry:

## Slide 4: Calibration Data
Maintained by SALT (commissioning data):
Linear Pol Zero-Point. Observe known unpolarized standards with several gratings
RSSpol_Linear_TelZeropoint_20061030_v01.txt
Linear Pol Efficiency, PA. Observe quartz lamp through polaroid with several gratings
RSSpol_HW_Calibration_20061030_v04.txt
Linear Pol PA Zero-Point. Observe known polarized standard
RSSpol_Linear_PAZeropoint.txt: 120

## Slide 5: Applying the Calibration
In specpolfinalstokes:
  for each rawstokes, zeropoint rotated from telescope to spectrograph frame and removed
  after combination of all rawstokes for the observation, polarization divided by HW efficiency
  PA rotated to spectrograph frame using HW PA data, then to sky using Slit PA and PA zeropoint
Fits header will contain:
  calibration file names
  slit PA, rotator rho, telescope azimuth

## Slide 6: Current calibration precision
On-axis: variable small PA "ripple" for high-SN, highly polarized objects
Off-axis (long slit): PA ripple + zeropoint change
 | Precision |  | 
Calibration | On-Axis | Full FOV | Full Track
Linear Pol Zero-Point | 0.05% | 1.5% | 0.1%?
Linear Pol Efficiency | 0.2 deg ripple | 0.5 deg ripple | 0.5 deg ripple
Linear Pol PA Zero-Point | 0.1 deg |  | 
PA ripple due to assembly alignment errors in 6 elements of halfwave plate
Variation due to sensitivity of waveplate to change in illumination of pupil over track and FOV (depends slightly on grating)

## Slide 7: Calibration Plans for future
(This year) Finish calibration revision to take into account illumination variations over track, FOV
  to be released in githib as new version of specpolfinalstokes, starting with PG0300, PG0900 gratings
(next year) Finish calibration revision for other gratings, new imaging spectropolarimetry mode
  to be released in github as new polsalt version, incorporating MOS capability, new wavecal and extraction
Do circular polarimetry calibration
Extend calibration into UV

## Slide 8: Questions? Extra added attractions?

## Slide 9: MOS, Imaging SpectroPolarimetry
MOS spectropolarimetry
  need zeropoint calibration over full FOV
Imaging spectropolarimetry
  dispersed vertically by beamsplitter chromatic splitting
    slitless. Find highly polarized objects
    MOS. very lores spectropolarimetry to photon limit
MOS support, needs
  mask design
  extraction
M30
20200822

## Slide 10: Circular Polarimetry
Calibration data needed
  Circular Pol Zero-Point?, likely negligible
  Circular Pol Efficiency (Waveplate)
  Linear-Circular Conversion (ADC, Waveplate)
Have started Linear-Circular Conversion by observing known star with high linear, zero circular
Have ongoing commissioning program to observe known magnetic stars for Zeeman Effect

## Slide 11: UV Polarimetry
UV spectropolarimetry very rare on large telescopes
Not currently feasible: arc optics is opaque below 3550 Ang
New calibration will be good to 3900 Ang (polaroid limitation)
UV calibration differences
  lamp too weak, use cluster of hot stars observed through UV polaroid, over variety of tracks
Have old UV spectropolarimetry (T Pyx, 2011) that is not yet publishable because we do not have a calibration